Shared Resources

Flow Cytometry Core Facility

Training Guide

Fred Hutchinson Cancer Research Center

Revised: December 2020

Table of Contents

- Overview of the Cytometer Software	3

- Cell Analysis Instruments - Startup and Shutdown

  - BD LSRFortessa X-50 SORP / BD FACSymphony A5 SORP	4

  - BD FACSCelesta Special Order Research Product	7

  - BD LSR II Special Order Research Product	9

  - BD FACSCanto II	11

- Cell Sorting Instruments - Startup and Shutdown

  - BD FACSAria II / BD Symphony S6 SORP 	13

  - Sony MA900 Multi-Application Cell Sorter	16

- Instrument Configuration Overviews	18

Overview of the Cytometer Software

- Creating a new experiment: Under the ‘Experiments’ drop down select new experiment. This will bring up previously saved templates. Select your template or under the ‘general’ tab find ‘New Experiment with blank sample tube.’ Experiments must be labeled with your Fred Hutch ID first and then the day’s date (mmddyr). For example, if your Fred Hutch is user1 and the date is January 1, 2020, name the experiment user1 010120. Following this, you can add any unique identifiers for that specific experiment. This ensures that the data will be archived correctly.

- Cytometer settings file (also known as instrument settings in older files): This file contains parameter voltage settings, threshold set point, and laser delay settings that are used to acquire data. Only the FSC voltage, SSC voltage, and the threshold are adjustable. All other settings are predetermined in the control file.

- Cytometer settings file (global vs tube-specific): The global cytometer settings file is located directly under the experiment name. After collection this file is automatically copied to each tube run and is saved with each tube.

- Parameter usage: Determine the detectors needed for each experiment and delete all parameters not being used in the experiment. This is located in the cytometer window under parameters.

- Creating a new experimental template file: Right click on experiment name and export experiment template. Save the template under your name tab and use a template name with any unique identifiers needed.

- Exporting data: Right click on experiment name and export FCS v. 3.0 files (while other formats can be exported always export using this format). Export this data to your HomeLink drive. All data in the flow core facility is archived monthly and older data is removed from the cytometer. It is your responsibility to obtain your data when after you complete your experiment.

For more information on using cytometer software, please refer to software manuals on the BD or Sony websites.

BD LSRFortessa X-50 SORP / BD FACSymphony A5 SORP: Startup and Shutdown

Filling the Sheath Tank and Emptying Waste Tank

- Check the weight of the sheath tank, a full tank (6L) is between 11 and 13 kg. Tanks should never run empty. A good indicator of the liquid level in the sheath tank is by checking how full the waste tank is.

- Empty waste tank: Unscrew the waste level sensor, remove it, and pour waste down the sink.

- Do not unscrew at the connection of the waste level sensor as this will break the sensor.

- Fill sheath tank: Pull pressure release valve (key ring) to release the pressure, unlatch air and sheath quick connects, and disconnect the sheath sensor. Remove lid and fill the sheath tank with sheath fluid (non-sterile 1X PBS) to the top weld line. Replace the lid and ensure that it is tight and secured. Reconnect the air and sheath quick connects and the sheath sensor.

- Do not unscrew at the connection of the sheath level sensor as this will break the sensor.

- Do not overfill the sheath tank as this will lead to back pressure and regulator issues

Starting Up Cytometer and Performing Fluidics Startup

- Start the cytometer by pressing the GREEN button on the right side of the instrument.

- Press the PRIME button.

- Once the prime is complete the instrument will reset, and the STANDBY button will illuminate.

- Place a FACS tube with ~ 2 mL 40% Contrad onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of blue detergent (Coulter Clenz) onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of 70% ethanol onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Bleed the sheath filter on the tank so that the line is free of air bubbles.

Starting Up Computer, Lasers, and Software

- Turn on the computer by pressing the computer power button.

- Log on to the ‘admin’ user for windows using the password: BDIS#1

- Open “Fiber Laser” software by clicking this button:

- If the ‘current,’ ‘power,’ and ‘SHG temp. values are blank without green numbering, then you are not connected to the correct com port (see figure to below). To connect to the right com port click this button: . Find the serial port where you see green number in the ‘current,’ ‘power,’ and ‘SHG temp. values.

  - Note: If none of serial ports are showing green numbering, restart the cytometer and computer.

- Type 200 into the power.

- Click On button.

- Click Activate button.

- The current should increase and the power should be ~200 mW.

- Open ‘BD Coherent’ icon on the desktop and the coherent connection software will open (see figure below)

- The 355 nm laser does not automatically start when opening the program and must be started manually. Enter 65 as the mW power then click ‘laser STOP’ and ‘laser START’. All lasers should be at the set power.

- Open BD FACSDIVA by clicking this button:

- Log in to Administrator profile by selecting Administrator in the drop down. No password is needed.

- The cytometer will connect to the computer.

  - Note: if the cytometer does not connect to the computer, restart the cytometer and computer.

Running Cytometer Setup and Tracking (CST) Beads

- Under ‘Cytometer’ select CST. This will open up a window for running a cytometer setup and tracking performance check.

- Load a FACS tube with ~2 mL of the appropriate CST beads in sterile 1x PBS.

- Match CST bead lot in FACS tube with the bead lot number set in the CST settings on the right of the CST window. If the bead lot is set to 00000, check with the flow cytometry core lab staff for the current bead lot number for this cytometer.

- Press RUN button and LOW button on cytometer.

- Ensure “Check Performance” is clicked. Click ‘Run’ on CST window.

- If performance check passes, click ‘finish’ and exit out of the CST setup window.

- If performance check does not pass, perform the following steps:

  - Ensure the cytometer is not clogged.

  - Prime cytometer.

  - Vortex beads or prepare new beads.

  - If the CST performance still does not pass, proceed to the Ultra Rainbow section below and check the CV values of each parameter. If they are all <10 then it is okay to run your samples.

- Remove FACS tube with CST beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Once the CST window is closed and the cytometer reconnects, a prompt will appear asking what settings you would like to use. You will select “Use CST Settings” as you have just run a performance check that passed on the CST beads.

Running Ultra-Rainbow (UR) 3.8 um Beads

- Open the file for the current year and select the experiment for current month. Expand out the specimen and add a new tube label this tube with the current day’s date: mmddyr

- Click green arrow to the left of the tube you named with today’s date.

- Load FACS tube with ~2 mL of UR beads.

- Press RUN and LOW and set fine adjust to 250.

- Click acquire and ensure that the events are within the set gates and click record to record 10,000 events.

- Remove FACS tube with UR beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Record UR bead values in log book.

Preparing to Run User Experiment

- In DIVA software, click File and Log Out.

- Log into FHCRC profile. No password is needed.

- Create a new daily folder and label it with the current day’s date: mmddyr

- Run your experiment.

Shutdown

- Place a FACS tube with ~ 2 mL 10% bleach (NaOCl) onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of blue detergent (Coulter Clenz) onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Press the GREEN button on the right side of the cytometer to turn off the machine.

- Exit out of BD FACSDIVA software.

- Click the windows home icon and select shutdown.

- Ensure that the cytometer and computer turned off.

BD FACSCelesta Special Order Research Product: Startup and Shutdown

Filling the Sheath Tank and Emptying Waste Tank

- Check the weight of the sheath tank, a full tank (6L) is between 11 and 13 kg. Tanks should never run empty. A good indicator of the liquid level in the sheath tank is by checking how full the waste tank is.

- Empty waste tank: Unscrew the waste level sensor, remove it, and pour waste down the sink.

- Do not unscrew at the connection of the waste level sensor as this will break the sensor.

- Fill sheath tank: Pull pressure release valve (key ring) to release the pressure, unlatch air and sheath quick connects, and disconnect the sheath sensor. Remove lid and fill the sheath tank with sheath fluid (non-sterile 1X PBS) to the top weld line. Replace the lid and ensure that it is tight and secured. Reconnect the air and sheath quick connects and the sheath sensor.

- Do not unscrew at the connection of the sheath level sensor as this will break the sensor.

- Do not overfill the sheath tank as this will lead to back pressure and regulator issues

Starting Up Cytometer and Performing Fluidics Startup

- Start the cytometer by pressing the GREEN button on the right side of the instrument.

- Press the PRIME button.

- Once the prime is complete the instrument will reset, and the STANDBY button will illuminate.

- Place a FACS tube with ~ 2 mL 40% Contrad onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of blue detergent (Coulter Clenz) onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of 70% ethanol onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Bleed the sheath filter on the tank so that the line is free of air bubbles.

Starting Up Computer, Lasers, and Software

- Turn on the computer by pressing the computer power button.

- Log on to the ‘admin’ user for windows using the password: BDIS#1

- The lasers will automatically turn on. Do not open coherent connect software.

- Open BD FACSDIVA by clicking this button:

- Log in to Administrator profile by selecting Administrator in the drop down. No password is needed.

- The cytometer will connect to the computer.

  - Note: if the cytometer does not connect to the computer, restart the cytometer and computer.

Running Cytometer Setup and Tracking (CST) Beads

- Under ‘Cytometer’ select CST. This will open up a window for running a cytometer setup and tracking performance check.

- Load a FACS tube with ~2 mL of the appropriate CST beads in sterile 1x PBS.

- Match CST bead lot in FACS tube with the bead lot number set in the CST settings on the right of the CST window. If the bead lot is set to 00000, check with the flow cytometry core lab staff for the current bead lot number for this cytometer.

- Press RUN button and LOW button on cytometer.

- Ensure “Check Performance” is clicked. Click ‘Run’ on CST window.

- If performance check passes, click ‘finish’ and exit out of the CST setup window.

- If performance check does not pass, perform the following steps:

  - Ensure the cytometer is not clogged.

  - Prime cytometer.

  - Vortex beads or prepare new beads.

  - If the CST performance still does not pass, proceed to the Ultra Rainbow section below and check the CV values of each parameter. If they are all <10 then it is okay to run your samples.

- Remove FACS tube with CST beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Once the CST window is closed and the cytometer reconnects, a prompt will appear asking what settings you would like to use. You will select “Use CST Settings” as you have just run a performance check that passed on the CST beads.

Running Ultra-Rainbow (UR) 3.8 um Beads

- Open the file for the current year and select the experiment for current month. Expand out the specimen and add a new tube label this tube with the current day’s date: mmddyr

- Click green arrow to the left of the tube you named with today’s date.

- Load FACS tube with ~2 mL of UR beads.

- Press RUN and LOW and set fine adjust to 250.

- Click acquire and ensure that the events are within the set gates and click record to record 10,000 events.

- Remove FACS tube with UR beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Record UR bead values in log book.

Preparing to Run User Experiment

- In DIVA software, click File and Log Out.

- Log into FHCRC profile. No password is needed.

- Create a new daily folder and label it with the current day’s date: mmddyr

- Run your experiment.

Shutdown

- Place a FACS tube with ~ 2 mL 10% bleach (NaOCl) onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of blue detergent (Coulter Clenz) onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and run on ‘high’ at a flow rate of 500 for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Press the GREEN button on the right side of the cytometer to turn off the machine.

- Exit out of BD FACSDIVA software.

- Click the windows home icon and select shutdown.

- Ensure that the cytometer and computer turned off.

# BD LSR II Special Order Research Product: Startup and Shutdown

Filling the Sheath Tank and Emptying Waste Tank

- Check the weight of the sheath tank, a full tank (6L) is between 11 and 13 kg. Tanks should never run empty. A good indicator of the liquid level in the sheath tank is by checking how full the waste tank is.

- Empty waste tank: Unscrew the waste level sensor, remove it, and pour waste down the sink.

- Do not unscrew at the connection of the waste level sensor as this will break the sensor.

- Fill sheath tank: Pull pressure release valve (key ring) to release the pressure, unlatch air and sheath quick connects, and disconnect the sheath sensor. Remove lid and fill the sheath tank with sheath fluid (non-sterile 1X PBS) to the top weld line. Replace the lid and ensure that it is tight and secured. Reconnect the air and sheath quick connects and the sheath sensor.

- Do not unscrew at the connection of the sheath level sensor as this will break the sensor.

- Do not overfill the sheath tank as this will lead to back pressure and regulator issues

Starting Up Cytometer and Performing Fluidics Startup

- Start the cytometer by pressing the GREEN button on the bottom front right of the instrument.

- Press the PRIME button.

- Once the prime is complete the instrument will reset, and the STANDBY button will illuminate.

- Place a FACS tube with ~ 2 mL 40% Contrad onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of blue detergent (Coulter Clenz) onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of 70% ethanol onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Bleed the sheath filter on the tank so that the line is free of air bubbles.

Starting Up Computer, Lasers, and Software

- Turn on the computer by pressing the computer power button.

- Log on to the ‘admin’ user for windows using the password: BDIS#1

- Open “Fiber Laser” software by clicking this button:

- If the ‘current,’ ‘power,’ and ‘SHG temp. values are blank without green numbering, then you are not connected to the correct com port (see figure to right). To connect to the right com port click this button: . Find the serial port where you see green number in the ‘current,’ ‘power,’ and ‘SHG temp. values.

  - Note: If none of serial ports are showing green numbering, restart the cytometer and computer.

- Type 200 into the power.

- Click On button.

- Click Activate button.

- The current should increase and the power should be ~200 mW.

- Open BD FACSDIVA by clicking this button:

- Log in to Administrator profile by selecting Administrator in the drop down. No password is needed.

- The cytometer will connect to the computer.

  - Note: if the cytometer does not connect to the computer, restart the cytometer and computer.

Running Cytometer Setup and Tracking (CST) Beads

- Under ‘Cytometer’ select CST. This will open up a window for running a cytometer setup and tracking performance check.

- Load a FACS tube with ~2 mL of the appropriate CST beads in sterile 1x PBS.

- Match CST bead lot in FACS tube with the bead lot number set in the CST settings on the right of the CST window. If the bead lot is set to 00000, check with the flow cytometry core lab staff for the current bead lot number for this cytometer.

- Press RUN button and LOW button on cytometer.

- Ensure “Check Performance” is clicked. Click ‘Run’ on CST window.

- If performance check passes, click ‘finish’ and exit out of the CST setup window.

- If performance check does not pass, perform the following steps:

  - Ensure the cytometer is not clogged.

  - Prime cytometer.

  - Vortex beads or prepare new beads.

  - If the CST performance still does not pass, proceed to the Ultra Rainbow section below and check the CV values of each parameter. If they are all <10 then it is okay to run your samples.

- Remove FACS tube with CST beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Once the CST window is closed and the cytometer reconnects, a prompt will appear asking what settings you would like to use. You will select “Use CST Settings” as you have just run a performance check that passed on the CST beads.

Running Ultra-Rainbow (UR) 3.8 um Beads

- Open the file for the current year and select the experiment for current month. Expand out the specimen and add a new tube label this tube with the current day’s date: mmddyr

- Click green arrow to the left of the tube you named with today’s date.

- Load FACS tube with ~2 mL of UR beads.

- Press RUN and LOW.

- Click acquire and ensure that the events are within the set gates and click record to record 10,000 events.

- Remove FACS tube with UR beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Record UR bead values in log book.

Preparing to Run User Experiment

- In DIVA software, click File and Log Out.

- Log into FHCRC profile. No password is needed.

- Create a new daily folder and label it with the current day’s date: mmddyr

- Run your experiment.

Shutdown

- Place a FACS tube with ~ 2 mL 10% bleach (NaOCl) onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of blue detergent (Coulter Clenz) onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and run on ‘high’ for 5 minutes.

- Remove FACS tube and press PRIME button.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Press the GREEN button on the bottom front right of the cytometer to turn off the machine.

- Exit out of BD FACSDIVA software.

- Click the windows home icon and select shutdown.

- Ensure that the cytometer and computer turned off.

BD FACSCanto II: Startup and Shutdown

Filling the Sheath Tank and Emptying Waste Tank

- Check the fluid levels. This can be done by checking their weight or opening the lids. The weight of the sheath tank when full (20L) is approximately 20 kg. Tanks should never run empty. There is a level indicator in the diva program on the ‘cytometer’ window.

- Empty waste tank: Unscrew the waste level sensor, remove it, and pour waste down the sink.

- Fill sheath tank: Disconnect the sheath sensor. Fill the sheath tank with sheath fluid (non-sterile 1X PBS). Replace the sheath sensor.

  - Do not overfill the sheath tank as this will lead to back pressure and regulator issues

- Bleed the three fluid filters on the front of the fluidics cart by unscrewing the bleed port. These are gravity fed and pressure from the tanks should force air out of the bleed ports. Replace the plugs once the air is removed.

Starting Up Cytometer, Computer, and Performing Fluidics Startup

- Start the cytometer by pressing the GREEN button on the left side of the instrument.

- Turn on the computer by pressing the computer power button.

- Log on to the ‘admin’ user for windows using the following password:

  - Canto 2 - 1 password: BDIS#1

  - Canto 2 - 2 password: BDIS

- The lasers will automatically turn on.

- Open BD FACSDIVA by clicking this button:

- Log in to Administrator profile by selecting Administrator in the drop down. No password is needed.

- The cytometer will connect to the computer.

  - Note: if the cytometer does not connect to the computer, restart the cytometer and computer.

- A prompt will then appear asking what settings you would like to use. You will select “Use CST Settings” as you have just run a performance check that passed on the CST beads.

- Attach HTS coupler.

- From the drop down, menu at the top select ‘Cytometer,’ then select ‘Cleaning Modes’ and ‘Prime Tank after Fill.’ A window will appear with three boxes that can be check marked. Check all three boxes and select ‘Run’

  - Perform this action twice.

- From the drop down, menu at the top select ‘Cytometer,’ then select ‘Bubble Filter Purge and Degas Flow Cell’.

  - Perform this action twice.

- From the drop down, menu at the top select ‘Cytometer,’ then select ‘Fluidics Startup’. A window will appear notifying you that this action will take 10 minutes to complete. Click ‘Run’ to proceed.

  - If HTS coupler is not attached, it will prompt you to attach the HTS coupler.

- When fluidics startup is complete, click ok on the popup window.

Running Cytometer Setup and Tracking (CST) Beads

- Under ‘Cytometer’ select CST. This will open up a window for running a cytometer setup and tracking performance check.

- Load a FACS tube with ~2 mL of the appropriate CST beads in sterile 1x PBS.

- Match CST bead lot in FACS tube with the bead lot number set in the CST settings on the right of the CST window. If the bead lot is set to 00000, check with the flow cytometry core lab staff for the current bead lot number for this cytometer.

- Ensure “Check Performance” is clicked. Click ‘Run’ on CST window.

- If performance check passes, click ‘finish’ and exit out of the CST setup window.

- If performance check does not pass, perform the following steps:

  - Ensure the cytometer is not clogged.

  - Prime cytometer.

  - Vortex beads or prepare new beads.

  - If the CST performance still does not pass, proceed to the Ultra Rainbow section below and check the CV values of each parameter. If they are all <10 then it is okay to run your samples.

- Remove FACS tube with CST beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Once the CST window is closed and the cytometer reconnects, a prompt will appear asking what settings you would like to use. You will select “Use CST Settings” as you have just run a performance check that passed on the CST beads.

Running Ultra-Rainbow (UR) 3.8 um Beads

- Open the file for the current year and select the experiment for current month. Expand out the specimen and add a new tube label this tube with the current day’s date: mmddyr

- Click green arrow to the left of the tube you named with today’s date.

- Load FACS tube with ~2 mL of UR beads.

- Set flow rate to LOW.

- Click acquire and ensure that the events are within the set gates and click record to record 10,000 events.

- Remove FACS tube with UR beads from the cytometer.

- Place a FACS tube with ~2 mL of distilled H2O onto the cytometer and press STANDBY button.

- Record UR bead values in log book.

Preparing to Run User Experiment

- In DIVA software, click File and Log Out.

- Log into FHCRC profile. No password is needed.

- Create a new daily folder and label it with the current day’s date: mmddyr

- Run your experiment.

Shutdown

- Attach HTS coupler.

- From the drop down, menu at the top select ‘Cytometer,’ then select ‘Fluidics Shutdown’. A window will appear notifying you that this action will take 5 minutes to complete. Click ‘Run’ to proceed.

  - If HTS coupler is not attached, it will prompt you to attach the HTS coupler.

- Once completed. Click ‘Ok’ Press the GREEN button on the side of the cytometer to turn off the machine.

- Press the GREEN button on the left side of the cytometer to turn off the machine.

- Exit out of BD FACSDIVA software.

- Click the windows home icon and select shutdown.

- Ensure that the cytometer and computer turned off.

BD FACSAria II and BD Symphony S6: Startup and Shutdown

Filling the Sheath Tank and Emptying Waste Tank

- Check the weight of the sheath tank, a full tank (6L) is between 11 and 13 kg. Tanks should never run empty. A good indicator of the liquid level in the sheath tank is by checking how full the waste tank is.

- Empty waste tank: Unscrew the waste level sensor, remove it, and pour waste down the sink.

  - Do not unscrew at the connection of the waste level sensor as this will break the sensor.

- Fill sheath tank: Pull pressure release valve (key ring) to release the pressure, unlatch air and sheath quick connects, and disconnect the sheath sensor. Attach sheath hose and quick connect vent. Do not remove lid and fill the sheath tank with sheath fluid (sterile 1X PBS). Reconnect the air and sheath quick connects and the sheath sensor.

  - Do not unscrew at the connection of the sheath level sensor as this will break the sensor.

  - Do not overfill the sheath tank as this will lead to back pressure and regulator issues

- Reattach all tanks to the cytometer.

## Starting up the Cytometer, Computer, Lasers, and Software

- Start the cytometer by pressing the GREEN button on the side of the machine.

- Turn on the computer and log on to the ‘admin’ user (Arias) or ‘BDadmin’ user (Symphony S6) for Windows using the following password:

  - Aria 2 - 2 and Symphony S6: BDIS#1

  - Aria 2 - 3: BDIS

- Open “Fiber Laser” software by clicking this button:

- If the ‘current,’ ‘power,’ and ‘SHG temp. values are blank without green numbering, then you are not connected to the correct com port (see figure below). To connect to the right com port click this button: . Find the serial port where you see green number in the ‘current,’ ‘power,’ and ‘SHG temp. values.

  - Note: If none of serial ports are showing green numbering, restart the cytometer and computer.

- Type 200 into the power.

- Click On button.

- Click Activate button.

- The current should increase and the power should be ~200 mW.

- For Aria 2 - 2 only: Open ‘BD Coherent’ icon on the desktop and the coherent connection software will open (see figure below).

  - The 355 nm laser does not automatically start when opening the program and must be started manually. Enter 65 as the mW power then click ‘laser STOP’ and ‘laser START’. All lasers should be at the set power.

- Open BD FACSDIVA by clicking this button:

- Log in to Administrator profile by selecting Administrator in the drop down. No password is needed.

- The cytometer will connect to the computer.

  - Note: if the cytometer does not connect to the computer, restart the cytometer and computer.

Performing Fluidics Startup

- Sonicate nozzle that will be used in experiment. After sonicating, spray the nozzle with 70% ethanol and use the air compressor or canned air to completely dry the nozzle.

- Remove the closed loop nozzle and black nozzle holster.

- Twist up the corner of a Kimwipe and place it into the space where the nozzle would normally go. Also, insert a Kimwipe below the nozzle location above the deflection plates.

- Load a 5 mL FACS tube with 40% Contrad on the loader.

- Under “cytometer” from the top bar, select ‘cleaning modes’ and then ‘clean flow cell.’ Repeat this until the tube is dry.

- Load a 5 mL FACS tube with 70% ethanol on the loader.

- Under “cytometer” from the top bar, select ‘cleaning modes’ and then ‘clean flow cell.’ Repeat this multiple times.

- Remove the Kimwipes. Replace the closed loop nozzle and the black nozzle holster.

- Under ‘cytometer’ click ‘fluidics startup.’ Click Ok on all prompts after completing each step. This will run for 5-10minutes.

- Bleed the sheath filter on the tank so that the line is free of air bubbles.

- Wipe off the deflector plates and soak up any droplets where the closed loop nozzle and black nozzle holster were using a Kimwipe.

- Insert the clean, sonicated nozzle into where the closed loop nozzle was. Replace the black nozzle holster.

- Turn on the stream by clicking the red X above the stream image window and let run for 10 minutes to allow for purging of any residual air or ethanol.

### Obtaining Purity Sample

- Fill the tube containing ~2.5 mL of LB broth with about 1.0 mL of PBS from the sort stream.

- Replace cap and leave in biosafety cabinet to incubate at room temperature for 7 days.

Running Cytometer Setup and Tracking (CST) Beads

- Under ‘Cytometer’ select CST. This will open up a window for running a cytometer setup and tracking performance check. Make sure the configuration is for the nozzle that is being used in the experiment.

- Load a FACS tube with ~2 mL of the appropriate CST beads in sterile 1x PBS onto loader.

  - Match CST bead lot in FACS tube with the bead lot number set in the CST settings on the right of the CST window. If the bead lot is set to 00000, check with the flow cytometry core lab staff for the current bead lot number for this cytometer.

- Ensure “Check Performance” is clicked. Click ‘Run’ on CST window.

- If performance check passes, click ‘finish’ and exit out of the CST setup window.

- If performance check does not pass, perform the following steps:

  - Ensure the cytometer is not clogged and stream profile looks okay.

  - Vortex beads or prepare new beads.

  - If the CST performance still does not pass, proceed to the Ultra Rainbow section below and check the CV values of each parameter. If they are all <10 then it is okay to run your samples.

- Remove FACS tube with CST beads from the cytometer.

- Once the CST window is closed and the cytometer reconnects, a prompt will appear asking what settings you would like to use. You will select “Use CST Settings” as you have just run a performance check that passed on the CST beads.

Running Ultra-Rainbow (UR) 3.8 um Beads

- Open the file for the current year and select the experiment for current month. Expand out the specimen and add a new tube label this tube with the current day’s date: mmddyr

- Click green arrow to the left of the tube you named with today’s date.

- Place FACS tube with ~2 mL of UR beads onto loader.

- Click load. The cytometer will automatically start acquiring data.

- Set flow rate to 1.0.

- Ensure that the events are within the set gates and click record to record 10,000 events.

- Unload FACS tube.

Running Accudrop Beads

- In DIVA software, click File and Log Out.

- Log into FHCRC profile. No password is needed.

- Open accudrop delay experiment.

- Load Accudrop beads onto loader.

- Click load. The cytometer will automatically start acquiring data.

- Adjust flow rate such that the event rate is 2,000 – 3,000 events/second.

- Open up the sort layout. Set device to 2 tube. Set precision to Fine Tune. Make sure that the correct accudrop gate is selected for the beads to be sorted into the left.

- Turn on the voltage.

- Turn on the optical filter.

- Start the sort. Click ‘cancel’ on the pop-up note for the waste drawer.

- Adjust the deflection to ensure accudrops are centered in the optical filter box on the left.

- Adjust “drop delay” to maximize the sort stream percentage or run the “auto delay” program to get the percentage as close to 100% as possible.

- Unload FACS tube.

Preparing to Run User Experiment

- Create a new daily folder and label it with the current day’s date: mmddyr

- Run your experiment.

## Shutdown

- Click ‘Cytometer’ and select ‘fluidics shutdown.’ Then follow the prompts and click ‘OK’ after each one is complete. Allow the sorter to finish the shutdown procedure.

  - Remove nozzle from flow cell.

  - Install the closed loop nozzle into the flow cell.

  - Connect the air and fluidics lines to the stainless-steel ethanol tank

    - Do NOT remove the sheave filter from the sheath tank.

    - Make sure the ethanol tank is full.

  - Release pressure from sheath tank.

  - Place a FACS tube with ~5 mL FACS Rinse as the “cleaning solution” onto the loading port.

- Press the GREEN button on the left of the cytometer to turn off the machine.

- Exit out of BD FACSDIVA software.

- Click the windows home icon and select shutdown.

- Ensure that the cytometer and computer turned off.

Sony MA900: Startup and Shutdown

Filling the Sheath Tank and Emptying Waste Tank

- Before powering on the system, check the levels of sheath fluid and waste in the tanks.

- Empty waste tank: Remove the tank from the fluidics cart (disconnecting the red waste line) and disposing of the liquid in the sink before securely reconnecting the tank to the system.

- Fill sheath tank: Attach sheath hose to female-female adaptor and attach the female adaptor to sheath line quick connect. Do not remove lid and fill the sheath tank with sheath fluid (sterile 1X PBS). Reconnect the air and sheath quick connects to the wet cart.

Cleaning Collection Chamber and Checking Air Compressor

- Open the instrument’s collection chamber and remove the black plates from the instrument by carefully unscrewing them. Spray the plates with 70% ethanol and dry with a Kimwipe.

- Spray a Kimwipe with 70% ethanol and wipe down all surfaces inside the collection chamber. Dry with a new Kimwipe

- Check that the air compressor is on and functioning.

## Starting up the Cytometer, Computer, Lasers, and Software

- If a user has logged out of windows, log into the ‘fcm’ user with the password ‘fcm’.

- Power on the cytometer using the Power/Standby button and wait until it goes into Standby.

  - Do not continue if it says processing, error, or initializing. Wait until it is in Standby.

- Open the MA900 Cell Sorter Software from the desktop.

- Login to the MA Cell Sorter software using ‘user1’ and ‘password1’.

- Proceed through the software’s prompts for Chip Detection and Chip Loading. Use a new 100 um sorting chip each day.

- Scan the QR code on the packaging using the camera on the monitor. Follow the commands in the software, hitting ‘Next’ prior to loading the chip.

- Open the cytometer lid using the black “Push open” button and install the chip vertically into its slot, ensuring that the ‘100 um’ text is upright and facing you. Push down gently on the chip and wait for it to load.

  - A used chip will appear: discard any old sorting chips, but retain the cleaning chip.

- Proceed through the startup commands by ensuring that the check boxes for all four lasers are checked (they may not populate automatically depending on the MA900 version).

- Select the Standard option and hit Next.

- The instrument will now do a pressure and fluidics check. Watch to see if the flow stream is stable between flushes.

  - Jittery/jumpy streams are a sign that the instrument may not pass this startup step successfully due to ethanol or air in the system. If so, perform bubble filter purges until the stream is stable.

- As instructed by the software, look inside the lit sample chamber. Check to ensure that back dripping is occurring by looking for droplets coming from the sample line.

- It should take ~10 minutes for this setup step to complete.

### Obtaining Purity Sample

- Fill the tube containing ~2.5 mL of LB broth with about 1.0 mL of PBS from the sort stream.

- Replace cap and leave in biosafety cabinet to incubate at room temperature for 7 days.

Performing Automatic Calibration

- Click the prompts to perform the automatic calibration.

- Prepare the automatic calibration beads:

  - The automatic calibration beads are stored at room temperature near the cytometers.

  - Mix the beads very well by vortexing.

  - The beads are used undiluted. Use ~2 mL of bead-containing solution in a 5mL FACS tube to calibrate.

- Once the beads are loaded in the proper tube holder, initiate the Auto Calibration by selecting Standard, hit Okay, load the tube of well-mixed, undiluted ~1mL beads on the instrument and hit Start. This step should take ~10 minutes.

- If the calibration fails:

  - Check that the deflection plates are attached correctly

  - Perform the Chip Debubble, Vent Air, and Sheath Filter Air Debubble functions

  - If it continually fails, perform a chip exchange. Follow the prompts to restart the system and perform automatic calibration again.

- When finished, hit Okay, remove the beads and check the droplet image. If the green light is lit solidly (not off, not blinking), the instrument is ready for use.

  - If the green light continues to blink, a prompt will show informing you to reperform the sort calibration. From the cytometer menu, click sort calibration and follow the prompts as you did in the initial start up.

## Shutdown

- Select “Software and Hardware Shutdown” from the cytometer menu in the shutdown section and follow the prompts in the shutdown wizard.

- Click ‘start’.

- Click ‘normal cleaning’.

- Load 10 mL bleach in a 15 mL conical tube into the sample chamber.

- Click ‘start’.

- Click ‘next’.

- Click ‘normal cleaning’ and ‘keep DI water inside sample probe’.

- Load 10 mL distilled H2O into the sample chamber.

- Click ‘start’.

- Click ‘next’.

- Click ‘ok’ to shutdown.

- Ensure that the cytometer has turned off. Leave the computer on.

Instrument Configuration Overview

Cell Analysis Instruments

- BD LSRFortessa X-50 Special Order Research Product (1)

  - Fortessa X-50: 5 lasers (355 nm, 406 nm, 488 nm, 552 nm, 628 nm), 28 color detection

- BD FACSymphony A5 Special Order Research Product (4)

  - Symphony 1: 5 lasers (355 nm, 405 nm, 488 nm, 552 nm, 628 nm), 28 color detection

  - Symphony 2: 5 lasers (355 nm, 406 nm, 488 nm, 552 nm, 628 nm), 28 color detection

  - Symphony 3: 5 lasers (355 nm, 406 nm, 488 nm, 552 nm, 628 nm), 28 color detection

  - Symphony 4: 5 lasers (355 nm, 406 nm, 488 nm, 552 nm, 628 nm), 28 color detection

- BD FACSCelesta Special Order Research Product (3)

  - Celesta 1: 4 lasers (405 nm, 488 nm, 561 nm, 637 nm), 16 color detection

  - Celesta 2: 4 lasers (405 nm, 488 nm, 561 nm, 637 nm), 16 color detection

  - Celesta 3: 4 lasers (405 nm, 488 nm, 561 nm, 637 nm), 16 color detection

- BD LSRII Special Order Research Product (1)

  - LSR 2 - 2: 4 lasers (405 nm, 488 nm, 532 nm, 642 nm), 18 color detection

- BD FACSCanto II (2)

  - Canto 2 - 1: 3 lasers (405 nm, 488 nm, 633 nm), 8 color detection

  - Canto 2 - 2: 3 lasers (405 nm, 488 nm, 633 nm), 8 color detection

Cell Sorting Instruments

- BD Symphony S6 Special Order Research Product (2)

  - Symphony S6 - 1: 5 lasers (355 nm, 404 nm, 488 nm, 552 nm, 628 nm), 29 color detection

  - Symphony S6 - 2: 5 lasers (355 nm, 402 nm, 488 nm, 552 nm, 628 nm), 29 color detection

- BD FACSAria II (2)

  - Aria 2 - 2: 5 lasers (355 nm, 404 nm, 488 nm, 561 nm, 627 nm), 18 color detection

  - Aria 2 - 3: 4 lasers (405 nm, 488 nm, 552 nm, 633 nm), 14 color detection

- Sony MA900 Multi-Application Cell Sorter (3)

  - Sony MA900 - 1: 4 Lasers (405nm, 488 nm, 561nm, 638 nm), 12 color detection

  - Sony MA900 - 2: 4 Lasers (405nm, 488 nm, 561nm, 638 nm), 12 color detection

  - Sony MA900 - 3: 4 Lasers (405nm, 488 nm, 561nm, 638 nm), 12 color detection

For more information on the configurations, please refer to our website at fredhutch.org.