RNeasy Tips:

- Make sure kit is not over 9 months olds

- RLT buffer and RW1 both contain guanidine salt which will precipitate out below RT (20-25°C) make sure they don’t have precipitates by warming to 50°C for a few minutes then cooling to RT. Guanidine is an RNase inhibitor so using it slightly warm should not be a degradation problem.

- Make sure centrifuge is at RT (not just previously used at 4°C)

- Add an extra RPE wash (does not contain guanidine salt) and incubate RPE washes for 3-5 minutes at RT before centrifugation

- Wash columns completely and take care not to carry over any previous reagent to the next step

________________________________________________________________________

Clean-up to improve the 260/230 ratio after extraction:

Add 500 ul of RPE buffer to the samples

Transfer it to the column

Incubate at RT for 5 min

Centrifuge

Repeat wash/incubation with RPE

Continue with kit instructions.

Notes:

-We recommend skipping RLT and RW1, it contains guanidine salt and can decrease 260/230, just use buffer RPE.

-If you have less than 1 ug you can follow the cleanup protocol in the RNeasy Micro kit or RNeasy Minelute Cleanup kit.

-If you have more than 1 ug you can use the cleanup protocol in the RNeasy Mini kit.

-You can increase the number of washes of buffer RPE, just make sure it has the right vol of ethanol added.

-Do not to use smaller volumes of elution buffer than what is recommended in each kit.